## Read me
KuBi — the equipment engine
Every asset carries its own record, and the record generates the work.
Generated from the engine. Regenerate with scripts/build-assets-xlsx.py.
The rule
"If Next Service Date = Today, the app automatically creates the service task." It does — and it also brings the task forward by the warning period, because a vendor cannot be booked on the morning the service falls due. What fires ON the day is the overdue state, which is louder.
The sheets
Sheet 'Your register' | START HERE. The blank for the clinic’s own asset list. Paste yours in, or send the Google Sheet and it gets loaded for you.
Sheet 'Register' | The 44 synthetic assets currently in the engine, for reference.
Sheet 'Cycles' | All 73 recurring obligations — what, how often, whose, and whether it stops the asset being used.
Sheet 'By role' | What each role carries across the whole register.
Sheet 'How it decides' | The four states an asset can be in, and why UNSERVICED is not the same as nothing being due.
Totals
Assets in the engine
Recurring obligations
That stop the asset being used
Assets whose failure stops the clinic
Rows filled into Your register
The register in this workbook is synthetic — the tags, serial numbers and contract dates are made up so the engine had something real to run on. Send the clinic’s own asset list and it replaces this one row for row, without a line of the engine changing.

## Your register
Your register — paste the clinic’s asset list here
Every column below is one the engine actually uses. The two in YELLOW HEADERS are the ones that make it work: without a last service date and an interval, no due date can be calculated — and the engine will report the asset as UNSERVICED rather than as nothing being due.
Asset ID | Name | Category | Location | Responsible role | Criticality | Make | Model | Serial | Commissioned | Warranty until | LAST SERVICE | SERVICE INTERVAL | Daily check — with its pass test | AMC vendor / expiry / cover
AUTOCLAVE-01 | Autoclave | Sterilization | Sterilisation room | Sterilization technician | Clinic stops | Runyes | Class B 23L | RY-B23-4471 | 2023-04-11 | 2025-04-10 | 2026-02-14 | Six-monthly | Water level, gasket intact, helix pack passes on the first cycle | Runyes India / 2026-11-14 / two visits a year, gasket and filter included
Note on C4: One of: Chairside, Sterilization, Imaging, Plant, Facility, Safety, Front office, Laboratory.
Note on E4: A ROLE, never a person — the record has to survive somebody leaving. One of: Dental assistant, Senior assistant, Sterilization technician, Reception, Housekeeping, Clinic manager, Lab coordinator, Inventory coordinator.
Note on F4: What breaks when it fails: "Clinic stops", "One room stops", or "Day continues".
Note on L4: The date it was ACTUALLY last serviced. Leave blank if genuinely unknown — the engine reports UNSERVICED, which is louder than overdue, rather than guessing.
Note on N4: Not "switch it on". A daily check needs a pass test, or it becomes a tick somebody does from the corridor.
Note on A5: An example row, showing the format. Overwrite it or delete it.

## Register
The 44 assets currently in the engine
Synthetic, and here for reference — the shape your own list will take.
Asset ID | Name | Category | Location | Responsible | Criticality | Make | Model | Cycles | Blocking | Daily check | AMC | Documents
CHAIR-01 | Dental chair unit 1 | Chairside | Operatory 1 | Dental assistant | One room stops | Confident | Elite | 2 | 0 | Back, up/down and headrest all move; light comes on; spittoon flushes | Confident Service | Installation certificate, User manual
XRAY-01 | Intraoral X-ray 1 | Imaging | Operatory 1 | Senior assistant | One room stops | Genoray | Port-X | 3 | 2 | Powers up, exposure indicator works, cone is undamaged | no contract | AERB registration, Radiation survey report
RVG-01 | RVG sensor 1 | Imaging | Operatory 1 | Dental assistant | One room stops | Carestream | RVG 5200 | 1 | 0 | Captures a test image; cable has no kinks at the strain relief | no contract | User manual
SCALER-01 | Ultrasonic scaler 1 | Chairside | Operatory 1 | Dental assistant | Day continues | Woodpecker | UDS-J | 1 | 0 | Activates, water flows through the tip, tip is not worn short | no contract | none held
CURE-01 | Curing light 1 | Chairside | Operatory 1 | Dental assistant | One room stops | Ivoclar | Bluephase | 1 | 1 | Charged, emits, light guide is clean and uncracked | no contract | User manual
CHAIR-02 | Dental chair unit 2 | Chairside | Operatory 2 | Dental assistant | One room stops | Confident | Elite | 2 | 0 | Back, up/down and headrest all move; light comes on; spittoon flushes | Confident Service | Installation certificate, User manual
XRAY-02 | Intraoral X-ray 2 | Imaging | Operatory 2 | Senior assistant | One room stops | Genoray | Port-X | 3 | 2 | Powers up, exposure indicator works, cone is undamaged | no contract | AERB registration, Radiation survey report
RVG-02 | RVG sensor 2 | Imaging | Operatory 2 | Dental assistant | One room stops | Carestream | RVG 5200 | 1 | 0 | Captures a test image; cable has no kinks at the strain relief | no contract | User manual
SCALER-02 | Ultrasonic scaler 2 | Chairside | Operatory 2 | Dental assistant | Day continues | Woodpecker | UDS-J | 1 | 0 | Activates, water flows through the tip, tip is not worn short | no contract | none held
CURE-02 | Curing light 2 | Chairside | Operatory 2 | Dental assistant | One room stops | Ivoclar | Bluephase | 1 | 1 | Charged, emits, light guide is clean and uncracked | no contract | User manual
CHAIR-03 | Dental chair unit 3 | Chairside | Operatory 3 | Dental assistant | One room stops | Confident | Elite | 2 | 0 | Back, up/down and headrest all move; light comes on; spittoon flushes | Confident Service | Installation certificate, User manual
XRAY-03 | Intraoral X-ray 3 | Imaging | Operatory 3 | Senior assistant | One room stops | Genoray | Port-X | 3 | 2 | Powers up, exposure indicator works, cone is undamaged | no contract | AERB registration, Radiation survey report
RVG-03 | RVG sensor 3 | Imaging | Operatory 3 | Dental assistant | One room stops | Carestream | RVG 5200 | 1 | 0 | Captures a test image; cable has no kinks at the strain relief | no contract | User manual
SCALER-03 | Ultrasonic scaler 3 | Chairside | Operatory 3 | Dental assistant | Day continues | Woodpecker | UDS-J | 1 | 0 | Activates, water flows through the tip, tip is not worn short | no contract | none held
CURE-03 | Curing light 3 | Chairside | Operatory 3 | Dental assistant | One room stops | Ivoclar | Bluephase | 1 | 1 | Charged, emits, light guide is clean and uncracked | no contract | User manual
CHAIR-04 | Dental chair unit 4 | Chairside | Operatory 4 | Dental assistant | One room stops | Confident | Elite | 2 | 0 | Back, up/down and headrest all move; light comes on; spittoon flushes | Confident Service | Installation certificate, User manual
XRAY-04 | Intraoral X-ray 4 | Imaging | Operatory 4 | Senior assistant | One room stops | Genoray | Port-X | 3 | 2 | Powers up, exposure indicator works, cone is undamaged | no contract | AERB registration, Radiation survey report
RVG-04 | RVG sensor 4 | Imaging | Operatory 4 | Dental assistant | One room stops | Carestream | RVG 5200 | 1 | 0 | Captures a test image; cable has no kinks at the strain relief | no contract | User manual
SCALER-04 | Ultrasonic scaler 4 | Chairside | Operatory 4 | Dental assistant | Day continues | Woodpecker | UDS-J | 1 | 0 | Activates, water flows through the tip, tip is not worn short | no contract | none held
CURE-04 | Curing light 4 | Chairside | Operatory 4 | Dental assistant | One room stops | Ivoclar | Bluephase | 1 | 1 | Charged, emits, light guide is clean and uncracked | no contract | User manual
AUTOCLAVE-01 | Autoclave | Sterilization | Sterilisation room | Sterilization technician | Clinic stops | Runyes | Class B 23L | 4 | 3 | Water level, door gasket intact, Bowie-Dick or helix pack passes on the first cycle | Runyes India | Installation qualification, User manual, Validation certificate
ULTRASONIC-01 | Ultrasonic cleaner | Sterilization | Sterilisation room | Sterilization technician | One room stops | Confident | CD-4820 | 3 | 1 | Solution level and freshness; cavitation confirmed on a foil test weekly | no contract | User manual
SEALER-01 | Pouch sealing machine | Sterilization | Sterilisation room | Sterilization technician | Clinic stops | Hawo | hm 500 DC-V | 2 | 1 | Seal temperature reached; a test pouch peels cleanly with no channels | no contract | none held
DISTILLER-01 | Water distiller | Sterilization | Sterilisation room | Sterilization technician | Clinic stops | Megahome | MH943 | 1 | 0 | Output conductivity acceptable; boiling chamber not scaled | no contract | none held
UV-01 | UV cabinet — sterilisation room | Sterilization | Sterilisation room | Sterilization technician | Day continues | Confident | UV-90 | 1 | 1 | UV lamp lights and the interlock stops it with the door open | no contract | none held
NEEDLE-01 | Needle destroyer | Sterilization | Sterilisation room | Sterilization technician | Day continues | Sharps | ND-2 | 1 | 0 | Burns through a test needle; collection chamber below the fill line | no contract | none held
COMPRESSOR-01 | Air compressor | Plant | Plant room | Clinic manager | Clinic stops | Josef Kaeser | Dental 3 | 4 | 2 | Switched on with a 5-minute warm-up; pressure gauge reads working range; tank drained | Airtech Services | Pressure vessel certificate, User manual
SUCTION-01 | Central suction pump | Plant | Plant room | Clinic manager | Clinic stops | Cattani | Turbo Smart | 3 | 2 | Suction strength at the furthest operatory; filters clear | Cattani Care — NO EXPIRY RECORDED | User manual
PUMP-01 | Water pump | Facility | Plant room | Housekeeping | Clinic stops | Crompton | Mini Champ | 1 | 0 | Started at opening; overhead tank filling; no dry running | no contract | none held
OPG-01 | OPG and cephalometric unit | Imaging | X-ray room | Senior assistant | One room stops | Vatech | PaX-i | 3 | 2 | Powers up, positioning lasers align, test exposure is clean | Vatech India | AERB licence, Radiation survey report, Installation certificate
SCANNER-01 | Intraoral scanner | Chairside | Mobile — trolley | Senior assistant | Day continues | Medit | i700 | 1 | 1 | Boots, tip heater reaches temperature, a test scan stitches | no contract | User manual, Calibration record
LAB-MICRO-01 | Laboratory micromotor | Laboratory | Laboratory | Senior assistant | Day continues | Marathon | N7 | 1 | 0 | Runs smoothly through its range with no vibration | no contract | none held
TRIMMER-01 | Model trimmer | Laboratory | Laboratory | Senior assistant | Day continues | Confident | MT-10 | 1 | 0 | Water flows over the disc; disc is not glazed | no contract | none held
UPS-01 | UPS and inverter | Facility | Plant room | Clinic manager | Clinic stops | Luminous | Zelio 1100 | 2 | 1 | On mains, battery indicator healthy, no alarm | no contract | none held
AC-RECEPTION | Air conditioner — reception | Facility | Reception | Housekeeping | Day continues | Daikin | FTKF50 | 1 | 0 | Set to 24 °C and reaching it; no water dripping from the indoor unit | CoolCare | none held
PURIFIER-01 | Air purifier and diffuser | Facility | Reception | Housekeeping | Day continues | Philips | AC2887 | 1 | 0 | Running; filter indicator not red | no contract | none held
RO-01 | RO water purifier | Facility | Pantry | Housekeeping | Day continues | Kent | Grand Plus | 1 | 0 | Dispenses; no leak at the tap | no contract | none held
FRIDGE-01 | Refrigerator — medicaments | Facility | Sterilisation room | Dental assistant | One room stops | Godrej | RD Edge | 1 | 1 | Temperature between 2 and 8 °C, read and written down | no contract | none held
EMERGENCY-01 | Emergency drug kit | Safety | Sterilisation room | Senior assistant | Clinic stops | — | Clinic kit | 2 | 1 | Kit present, sealed, and nothing expiring inside 30 days | no contract | Contents list with expiry dates
OXYGEN-01 | Oxygen cylinder and mask | Safety | Sterilisation room | Senior assistant | Clinic stops | INOX | D-type | 2 | 2 | Cylinder pressure above the refill line; mask and tubing present and sealed | no contract | Cylinder test certificate
FIRE-01 | Fire extinguisher — CO2 | Safety | Corridor | Clinic manager | Day continues | Ceasefire | CO2 4.5kg | 1 | 1 | none | no contract | Refill certificate
DVR-01 | CCTV recorder | Front Office | Reception | Reception | Day continues | Hikvision | DS-7108 | 1 | 0 | Recording, all cameras live, disk not full | no contract | none held
SERVER-01 | Practice computer and backup | Front Office | Reception | Clinic manager | Clinic stops | Dell | OptiPlex 7010 | 2 | 1 | Backup ran last night and completed | no contract | none held
CARDPOS-01 | Card machine | Front Office | Reception | Reception | Day continues | Pine Labs | A910 | 1 | 0 | Powers up, connects, roll loaded | no contract | none held
By category
Chairside |  | 13
Facility |  | 6
Front Office |  | 3
Imaging |  | 9
Laboratory |  | 2
Plant |  | 2
Safety |  | 3
Sterilization |  | 6

## Cycles
The 73 recurring obligations
Each one is its own clock. An autoclave has four running at once — the service, the weekly spore test, the annual validation and the gasket — and they are not one date. Filter column G for the ones that stop an asset being used.
Asset ID | Asset | What is due | How often | Warn ahead | Whose | Blocks? | Why it matters | YOUR CORRECTION
CHAIR-01 | Dental chair unit 1 | Service due | Six-monthly | 14 days | Clinic manager |  | The service interval for this asset has come round
CHAIR-01 | Dental chair unit 1 | Upholstery inspection | Quarterly | 7 days | Dental assistant |  | Torn upholstery cannot be disinfected, and this one has not been looked at
XRAY-01 | Intraoral X-ray 1 | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
XRAY-01 | Intraoral X-ray 1 | AERB licence renewal | Two-yearly | 60 days | Clinic manager | YES | An X-ray unit operating on a lapsed AERB licence is operating illegally
XRAY-01 | Intraoral X-ray 1 | Radiation output and leakage check | Yearly | 30 days | Clinic manager | YES | The annual radiation quality assurance check is due
RVG-01 | RVG sensor 1 | Sensor cable inspection | Monthly | 3 days | Dental assistant |  | RVG sensors die at the cable, and this one has not been looked at
SCALER-01 | Ultrasonic scaler 1 | Scaler tip wear check | Quarterly | 7 days | Dental assistant |  | A scaler tip worn 2 mm short has lost half its power, and nobody has measured these
CURE-01 | Curing light 1 | Output measured on a radiometer | Monthly | 3 days | Dental assistant | YES | A curing light losing output cures nothing properly and looks identical, and this one has not been measured
CHAIR-02 | Dental chair unit 2 | Service due | Six-monthly | 14 days | Clinic manager |  | The service interval for this asset has come round
CHAIR-02 | Dental chair unit 2 | Upholstery inspection | Quarterly | 7 days | Dental assistant |  | Torn upholstery cannot be disinfected, and this one has not been looked at
XRAY-02 | Intraoral X-ray 2 | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
XRAY-02 | Intraoral X-ray 2 | AERB licence renewal | Two-yearly | 60 days | Clinic manager | YES | An X-ray unit operating on a lapsed AERB licence is operating illegally
XRAY-02 | Intraoral X-ray 2 | Radiation output and leakage check | Yearly | 30 days | Clinic manager | YES | The annual radiation quality assurance check is due
RVG-02 | RVG sensor 2 | Sensor cable inspection | Monthly | 3 days | Dental assistant |  | RVG sensors die at the cable, and this one has not been looked at
SCALER-02 | Ultrasonic scaler 2 | Scaler tip wear check | Quarterly | 7 days | Dental assistant |  | A scaler tip worn 2 mm short has lost half its power, and nobody has measured these
CURE-02 | Curing light 2 | Output measured on a radiometer | Monthly | 3 days | Dental assistant | YES | A curing light losing output cures nothing properly and looks identical, and this one has not been measured
CHAIR-03 | Dental chair unit 3 | Service due | Six-monthly | 14 days | Clinic manager |  | The service interval for this asset has come round
CHAIR-03 | Dental chair unit 3 | Upholstery inspection | Quarterly | 7 days | Dental assistant |  | Torn upholstery cannot be disinfected, and this one has not been looked at
XRAY-03 | Intraoral X-ray 3 | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
XRAY-03 | Intraoral X-ray 3 | AERB licence renewal | Two-yearly | 60 days | Clinic manager | YES | An X-ray unit operating on a lapsed AERB licence is operating illegally
XRAY-03 | Intraoral X-ray 3 | Radiation output and leakage check | Yearly | 30 days | Clinic manager | YES | The annual radiation quality assurance check is due
RVG-03 | RVG sensor 3 | Sensor cable inspection | Monthly | 3 days | Dental assistant |  | RVG sensors die at the cable, and this one has not been looked at
SCALER-03 | Ultrasonic scaler 3 | Scaler tip wear check | Quarterly | 7 days | Dental assistant |  | A scaler tip worn 2 mm short has lost half its power, and nobody has measured these
CURE-03 | Curing light 3 | Output measured on a radiometer | Monthly | 3 days | Dental assistant | YES | A curing light losing output cures nothing properly and looks identical, and this one has not been measured
CHAIR-04 | Dental chair unit 4 | Service due | Six-monthly | 14 days | Clinic manager |  | The service interval for this asset has come round
CHAIR-04 | Dental chair unit 4 | Upholstery inspection | Quarterly | 7 days | Dental assistant |  | Torn upholstery cannot be disinfected, and this one has not been looked at
XRAY-04 | Intraoral X-ray 4 | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
XRAY-04 | Intraoral X-ray 4 | AERB licence renewal | Two-yearly | 60 days | Clinic manager | YES | An X-ray unit operating on a lapsed AERB licence is operating illegally
XRAY-04 | Intraoral X-ray 4 | Radiation output and leakage check | Yearly | 30 days | Clinic manager | YES | The annual radiation quality assurance check is due
RVG-04 | RVG sensor 4 | Sensor cable inspection | Monthly | 3 days | Dental assistant |  | RVG sensors die at the cable, and this one has not been looked at
SCALER-04 | Ultrasonic scaler 4 | Scaler tip wear check | Quarterly | 7 days | Dental assistant |  | A scaler tip worn 2 mm short has lost half its power, and nobody has measured these
CURE-04 | Curing light 4 | Output measured on a radiometer | Monthly | 3 days | Dental assistant | YES | A curing light losing output cures nothing properly and looks identical, and this one has not been measured
AUTOCLAVE-01 | Autoclave | Service due | Six-monthly | 14 days | Clinic manager | YES | The service interval for this asset has come round
AUTOCLAVE-01 | Autoclave | Biological spore test | Weekly | 1 days | Sterilization technician | YES | A week has passed with no spore test, so nothing since the last one has been proven sterile
AUTOCLAVE-01 | Autoclave | Annual validation | Yearly | 45 days | Clinic manager | YES | The autoclave’s annual validation is due, and it cannot be booked on the day
AUTOCLAVE-01 | Autoclave | Door gasket replacement | Yearly | 21 days | Sterilization technician |  | The door gasket is a consumable and this one is at the end of its year
ULTRASONIC-01 | Ultrasonic cleaner | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
ULTRASONIC-01 | Ultrasonic cleaner | Foil cavitation test | Weekly | 1 days | Sterilization technician | YES | An ultrasonic cleaner that has stopped cavitating looks and sounds exactly the same
ULTRASONIC-01 | Ultrasonic cleaner | Solution changed | Daily | same day | Sterilization technician |  | The solution has not been changed today
SEALER-01 | Pouch sealing machine | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
SEALER-01 | Pouch sealing machine | Seal integrity test | Monthly | 3 days | Sterilization technician | YES | A failing seal is invisible until a pack is opened contaminated
DISTILLER-01 | Water distiller | Descale | Monthly | 3 days | Sterilization technician |  | A scaled distiller feeds the autoclave hard water, which is how chambers are ruined
UV-01 | UV cabinet — sterilisation room | UV lamp replacement | Yearly | 21 days | Sterilization technician | YES | A UV lamp past its hours still glows and no longer disinfects
NEEDLE-01 | Needle destroyer | Chamber emptied to the sharps stream | Weekly | 1 days | Sterilization technician |  | The needle destroyer chamber has not been emptied this week
COMPRESSOR-01 | Air compressor | Service due | Six-monthly | 14 days | Clinic manager | YES | The service interval for this asset has come round
COMPRESSOR-01 | Air compressor | Receiver drained of condensate | Daily | same day | Dental assistant |  | The compressor receiver has not been drained today, and standing water rusts it from the inside
COMPRESSOR-01 | Air compressor | Air filter changed | Six-monthly | 14 days | Clinic manager |  | The compressor air filter is due
COMPRESSOR-01 | Air compressor | Pressure vessel inspection | Two-yearly | 60 days | Clinic manager | YES | A pressure vessel inspection is a statutory certificate, not a service
SUCTION-01 | Central suction pump | Service due | Six-monthly | 14 days | Clinic manager | YES | The service interval for this asset has come round
SUCTION-01 | Central suction pump | Suction lines flushed and disinfected | Daily | same day | Dental assistant |  | The suction lines have not been flushed today
SUCTION-01 | Central suction pump | Amalgam separator emptied and disposed to record | Six-monthly | 21 days | Clinic manager | YES | An amalgam separator past capacity discharges mercury to drain
PUMP-01 | Water pump | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
OPG-01 | OPG and cephalometric unit | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
OPG-01 | OPG and cephalometric unit | AERB licence renewal | Two-yearly | 60 days | Clinic manager | YES | An OPG operating on a lapsed AERB licence is operating illegally
OPG-01 | OPG and cephalometric unit | Radiation output and image quality check | Yearly | 30 days | Clinic manager | YES | The annual radiation quality assurance check is due
SCANNER-01 | Intraoral scanner | Scanner calibration | Quarterly | 7 days | Senior assistant | YES | An uncalibrated scanner produces work the laboratory cannot fit
LAB-MICRO-01 | Laboratory micromotor | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
TRIMMER-01 | Model trimmer | Trimmer disc replacement | Yearly | 21 days | Senior assistant |  | A glazed trimmer disc burns models rather than cutting them
UPS-01 | UPS and inverter | Battery load test | Quarterly | 7 days | Clinic manager | YES | A UPS battery that has never been load tested is a UPS nobody has proved
UPS-01 | UPS and inverter | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
AC-RECEPTION | Air conditioner — reception | Filter clean | Quarterly | 7 days | Housekeeping |  | A dirty AC filter in a clinical space is an air quality problem, not a comfort one
PURIFIER-01 | Air purifier and diffuser | HEPA filter replacement | Six-monthly | 14 days | Housekeeping |  | A saturated HEPA filter is a fan
RO-01 | RO water purifier | Filter and membrane service | Six-monthly | 14 days | Housekeeping |  | The RO filters are past their interval, and this is the water that goes into patients’ mouths
FRIDGE-01 | Refrigerator — medicaments | Temperature log reviewed | Weekly | 1 days | Dental assistant | YES | A week of temperature readings nobody has looked at is a week of readings
EMERGENCY-01 | Emergency drug kit | Expiry check against the contents list | Monthly | 7 days | Senior assistant | YES | An emergency drug nobody has checked is an emergency drug that has expired
EMERGENCY-01 | Emergency drug kit | Medical emergency drill | Six-monthly | 21 days | Clinic manager |  | A kit nobody has practised with is a box
OXYGEN-01 | Oxygen cylinder and mask | Pressure checked and refill arranged | Monthly | 7 days | Senior assistant | YES | An oxygen cylinder nobody has read the gauge on is an oxygen cylinder of unknown volume
OXYGEN-01 | Oxygen cylinder and mask | Cylinder hydrostatic test | Five-yearly | 90 days | Clinic manager | YES | A cylinder past its hydrostatic test date may not legally be refilled
FIRE-01 | Fire extinguisher — CO2 | Extinguisher pressure test and refill | Yearly | 30 days | Clinic manager | YES | An extinguisher past its refill date is a red cylinder
DVR-01 | CCTV recorder | Recording retention confirmed | Monthly | 3 days | Reception |  | A DVR quietly overwriting at seven days is a DVR nobody can produce footage from
SERVER-01 | Practice computer and backup | Backup restore actually tested | Quarterly | 14 days | Clinic manager | YES | A backup nobody has restored from is a backup nobody has
SERVER-01 | Practice computer and backup | Service due | Yearly | 14 days | Clinic manager |  | The service interval for this asset has come round
CARDPOS-01 | Card machine | Batch settlement reconciled to the bank | Daily | same day | Reception |  | Today’s card batch has not been reconciled

## By role
What each role carries
The responsible person is held as a role, so the record survives somebody leaving.
Role | Assets | Cycles | Blocking | The ones that block, by name
Clinic manager | 5 | 37 | 20 | AERB licence renewal (OPG-01) · AERB licence renewal (XRAY-01) · AERB licence renewal (XRAY-02) · AERB licence renewal (XRAY-03) · AERB licence renewal (XRAY-04) · Amalgam separator emptied and disposed to record (SUCTION-01) · Annual validation (AUTOCLAVE-01) · Backup restore actually tested (SERVER-01) · Battery load test (UPS-01) · Cylinder hydrostatic test (OXYGEN-01) …
Dental assistant | 17 | 19 | 5 | Output measured on a radiometer (CURE-01) · Output measured on a radiometer (CURE-02) · Output measured on a radiometer (CURE-03) · Output measured on a radiometer (CURE-04) · Temperature log reviewed (FRIDGE-01)
Sterilization technician | 6 | 8 | 4 | Biological spore test (AUTOCLAVE-01) · Foil cavitation test (ULTRASONIC-01) · Seal integrity test (SEALER-01) · UV lamp replacement (UV-01)
Senior assistant | 10 | 4 | 3 | Expiry check against the contents list (EMERGENCY-01) · Pressure checked and refill arranged (OXYGEN-01) · Scanner calibration (SCANNER-01)
Housekeeping | 4 | 3 | 0 | —
Reception | 2 | 2 | 0 | —

## How it decides
The four states, and why the order matters
State | What it means | What the engine does
DOWN | Reported broken and not yet reported back. Downtime counts from the failure to now — it is the absence of a restoration event, not a status field somebody sets. | Shows first. Everything else waits.
UNSERVICED | Nobody has ever recorded a service against it. NOT the same as nothing being due: the arithmetic "next = last + interval" cannot run at all, and the obvious implementation quietly reports no service due — the exact opposite of the truth. | Ranked ABOVE overdue. A machine with no history is a machine nobody can vouch for.
OVERDUE | Serviced at some point, and the next date has passed. | The task exists with a day count against it.
OPERATIONAL | Working, and everything due against it is in date. | Quiet. Nothing is shown that nobody needs to act on.
Three things a checklist would show as one
No service interval at all · an interval nobody has ever met · an interval that has lapsed. A curing light has a monthly radiometer reading and no service contract; a fridge has a temperature log and no engineer. Reporting either as "never serviced" reads as neglect rather than as a machine that does not need servicing — so the engine holds the three apart and the screen shows them differently.
Warning ahead, not on the day
"Next Service Date = Today" is the floor, not the ceiling. An AMC that expires today cannot be renewed today, and an autoclave validation booked on the morning it falls due means a day with no sterilisation. So every cycle carries a warning period as well as an interval, and the work appears that many days early. What fires ON the day is the overdue state, which is a different and louder thing.